Overview and rationale on the use of substances in Medical Devices as controlled by
Directive 2002/95/EC on the restriction of the use of certain hazardous substances in electrical and electronic equipment (RoHS)

Explanation of each column

Modality

The generic name of the Medical device.

Substance

Usage of any restricted substance(s) listed in the RoHS directive.

Typical Application and Approximate amount per product

Function or part of a medical device in which the substance is applied and the amount of each Substance used.

Total Weight put on the EU market in kg/Year

Information in this column is derived from approximate units of Modalities marketed in the EU per year supplied by all manufacturers worldwide and the amount of Substance used by each Modality.

Benefit

Benefits, including medical, for using the Substance(s) in a Modality.

Replace

Lists any viable alternative to the restricted Substance. When considering this both technical alternatives and their associated financial implications are considered.

Rationale

This column gives the rationale for the continued use of a Substance (a) as listed in the RoHS directive by a Modality.

Colour coding in the table

|  | Lead |
| --- | --- |
|  | Cadmium |
|  | Hexavalent Chromium |
|  | Mercury |
|  | PBB, PBDE |
|  | To be completed weights |
| Text | To be discussed or corrected |
| Text | Added from member information |

| Modality | Sub-stance | Typical Application [Approx amount per product] |  | Total Weight put on the EU market in Kg | Benefit | Replace | Rationale |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| General Comment: The rationales given below can only take into account technology known to exist at the moment. Future technological developments may render some of the arguments obsolete, and this will need to be addressed at future reviews. For instance for some of the substance applications the industry proposes a review in 2012. |  |  |  |  |  |  |  |  |
| Radiology |  |  |  |  |  |  |  |  |
| Convent-ional X-Ray | Lead | Shielding X-Ray Tube/ Collimator [25kg] |  | 150000 | Reduction in unwanted X-Ray exposure | N | The effectiveness of x-ray shielding is a function of the density of the material. Any very dense material (e.g. Uranium) will shield against x-ray but may have other limitations. Lead offers the only realistic compromise between mass and volume. |  |
|  |  | Shielding Collimator, Raster [5-10kg] |  | 65000 | X-Ray dose reduction for patient/image quality | N | Thin strips of lead are used in x-ray beam path to reduce the amount of x-ray scatter. This enables lower x-ray doses to be given to the patient and also reduces film fogging enhancing image quality. |  |
|  |  | Electrical connections (eg PCB, connectors, wiring looms) [1kg] |  | 6500 | Product reliability (Patient safety, esp where the device controls the administration of energy and/or drugs to a patient) | Y, To be reviewed in 2012 | Lead solder is a well understood technology and consequently process reliability issues are to a large extent solved. Lead free solder is a relatively unknown process and so process reliability is still being developed. Other electrical Industries are converting to lead free and when the technology has stabilized, medical devices can make the change. |  |
|  |  | X-Ray tube rotating anode bearings [g] |  | 1 | Bearing life | N | X-ray tubes use rotating anodes to handle the large amount of heat generated during an x-ray exposure. Anode bearings are therefore subjected to high temperatures beyond the range of normal lubricants, also they are sealed for life within the tube evacuated glass envelope. Routine lubrication is therefore difficult. |  |
|  |  | Shielding Radiation shield X-Ray Image Intensifier [5-10kg] Radiation shields of other image detectors [7..17kg] |  | 47000 | Reduction in unwanted X-Ray exposure | N | Heavy metal such as tungsten alloy is difficult to machine work, difficult to recycle, expensive (approximately 10 times higher cost). In case of iron there is a disadvantage that the size and weight of equipment also becomes large.(see special issue on the replacement of lead) In general, the X-ray shielding can be achieved by adding the heavy metal in the glass component, but it must be achieved along with the optical capacity so that the use of Pb becomes essential. |  |
|  |  | Shielding (Well Type) Ionization Chambers [1-5kg] |  | 1000 | Improvement of precision by shielding from unwanted external irradiation | N | The effectiveness of x-ray shielding is a function of the density of the material. Any very dense material (e.g. Uranium) will shield against x-ray but may have other limitations. Lead offers the only realistic compromise between mass and volume. See special issue on the replacement of lead. |  |
|  |  | Shielding (Pediatric) X-ray Filters [5-100g] |  | 1000 | Reduction of radiation load especially to children | N | The effectiveness of x-ray shielding is a function of the density of the material. Lead oxide offers the possibility to provide transparent radiation filters. |  |
|  |  | X-ray test objects [1-10g] |  | 100 | Improve quality assurance | N cost | Difficult to replace by any other material at appropriate cost |  |
|  |  | Lead Acid Batteries [1-10 kg] AV to check lead acid directive |  | 15000 | - Mobile equipment moving power. - Power backup in case of power failure. - Examinations can be performed at sites where no power facility is available. | Y, But costly | To provide drive electric power of sufficiently large capacity and to keep safe and steady movement. Replacement is very costly. |  |
|  |  | Counterweights [50-100kg] |  | 80.000 | Counterbalance to movable heavy components to enable precise equipment positioning without patient involvement. | Y, In the future design | Some parts of an X-Ray set are heavy but have to move in order to take optimum pictures, as it may be impossible/undesirable to move the patient. Conventionally counterbalances consisted of lead to be effective with the lowest volume. Revised design solutions, and alternative approaches means lead counterweights can be eliminated over the coming years. |  |
|  | Lead+Cadmium (PbO,PbI2, CdTe,CdZn) | Radiation detectors [10-100g] |  | (100kg) | Reduction in unwanted X-Ray exposure Improvement to image quality What is used for shielding and what is used for detection? (BvdE) | N | FPI(Flat Panel Imager) is regarded as the device taking over existing X-ray Image Intensifier electron tube used for medical imaging or none-destructive inspection, because of its characteristic having less distortion of output image .One of the types of commercially available FPI is called indirect conversion type FPI, in which crystal film of cesium iodide with thallium additives or gadolinium oxide with terbium additives converts incident X-ray into visible light, and then photo diode array optically connected to the crystal film convert the visible light into electric charge to be read out by external electronic circuit. Because of the intrinsic property of visible light, it is inevitable that the light is scattered in the crystal film and that may result in blurred image of view. On the other hand Lead+ Cadmium type of FPI is called direct conversion type FPI, in which X-ray is directly converted into electric charge by photoconductive material film. There is no substitute for Lead+ Cadmium such high resolution &high conversion efficiency. |  |
|  | Cadmium | High Efficiency Phosphors [mg] and radiation detectors [mg] |  | 3 | Reduction in unwanted X-Ray exposure Improvement to image quality (BvdE) Used in Image Intensifier input and output screens | ## N, ## in future designs it is replaceable | It is impracticable to replace current method of producing the fluorescent substance. Disadvantage in health is expected due to the failure in reducing X-ray exposure to patient and manipulator. Input screen - Need for high X-ray absorption and high conversation efficiency Output screen – high contrast for small details |  |
|  | Cadmium(CdTe,CZT) | Radiation Semiconductor detector |  |  |  |  |  |  |
|  | Hexavalent chromium | Alkali dispensers of Photo cathode tubes |  | 3 | (BvdE) | N | To take them out of the tube would necessitate a major redesign of the production process and the product. |  |
|  | Mercury | Position Switches [g] |  | 1 |  | Y |  |  |
|  |  | Passive Safety Circuitry [g] |  | ? |  |  |  |  |
| Computed Tomography (CT) | Lead | Shielding X-Ray Tube/ Collimator [15kg] |  | 23000 | Reduction in unwanted X-Ray exposure | N | The effectiveness of x-ray shielding is a function of the density of the material. Any very dense material (e.g. Uranium) will shield against x-ray but may have other limitations. Lead offers the only realistic compromise between mass and volume. |  |
|  |  | Shielding Collimator [5kg] |  | 8000 | X-Ray dose reduction for patient and improvement of image quality | N | The collimator restricts the x-ray beam to the minimum necessary and thus enables lower x-ray doses to be given to the patient. It also reduces detector ‘noise’ enhancing image quality. |  |
|  |  | Shielding Shielding of the detector environment Raster collimator at the detector side […FS] |  |  | No radiation should escape from the CT scanner to the outside world |  |  |  |
|  |  | Counterweights In the rotating gantry [30 kg] |  | 12000 | Counterbalance to movable heavy components to enable precise equipment positioning without patient involvement. | Y, Future design | Some parts of an X-Ray set are heavy but have to move in order to take optimum pictures, as it may be impossible/undesirable to move the patient. Conventionally counterbalances have been lead to be effective with the lowest volume. Revised design solutions, and alternative approaches means lead counterweights can be eliminated over the next few years. |  |
|  | Cadmium | Collimators [kg] (PL) |  | 0 |  |  |  |  |
|  |  | High Efficiency Phosphors [5-15mg] (PL) |  | 0 |  |  |  |  |
|  |  | X-Ray detector crystals [400g] |  | 500 | Efficient and fast decay detector, reduction of radiation to the patients, improvement to image quality | N | It is impracticable to replace current detector material with the same benefits under the current state of knowledge. |  |
|  | Cadmium(CdTe,CZT) | Radiation Semiconductor detector |  |  |  |  |  |  |
|  | Mercury | X-Ray detector Crystals [g] |  | 0 | More efficient than Cd based detectors |  | Current research for future products Flexibility for innovations! |  |
|  | PBBE | Printed Circuit Boards |  |  | Reduction in flammability of PCBs? | Y, | The Medical Device industry uses PCB’s from the electronic industry which has to comply to the RoHS |  |
| Magnetic Resonance Imaging (MRI) | Lead | Supercooled magnet electrical contacts [7kg] |  | 2000 | To ensure the contact and improve magnet reliability | Y, To be reviewed in 2012 | Once the magnetic field is established, the magnet is effectively ‘shorted’ out and the established current flows in a supercooled environment (-196°C). The shorting contacts must therefore operate with lowest resistance at this temperature. The material (Wood’s Material) uses lead as one of its components. |  |
|  |  | Magnetic cooling materials [500g] (BvdE) |  | 50kg | Cooling for maintaining superconductivity | ### N | Difficult to cooling for maintaining superconductivity |  |
|  |  | Lead Acid Batteries[1-10 kg] |  |  | - Power backup in case of power failure. - Examinations can be performed at sites where no power facility is available. | Y, But costly | To provide drive electric power of sufficiently large capacity and to keep safe and steady movement. Replacement is very costly. |  |
|  | Cadmium | Supercooled magnet electrical contacts [2.5kg] |  | (7) | To ensure the contact and improve magnet reliability | Y, To be reviewed in 2012 | Once the magnetic field is established, the magnet is effectively ‘shorted’ out and the established current flows in a supercooled environment (-196°C). The shorting contacts must therefore operate with lowest resistance at this temperature. The material (Wood’s Material) uses cadmium as one of its components. |  |
|  | Hexavalent Chromium | Alloy material |  |  | Improve material properties | Y, in future design | E.g. in steel alloys. Used to improve strength properties. |  |
|  |  | Cooling system (FS, GD, PL) |  | 0 | Corrosion resistance |  |  |  |
| Convent-ional Nuclear Medicine, Single Photon Emission Computed Tomography (SPECT)and/or Positron Emission Tomography (PET) | Lead |  |  |  |  |  |  |  |
|  |  | Shielding Raster (collimator [15..300kg of lead] (FS, PL) |  | 100000 | Improved image quality, decrease of dose to the patient | N | Lead honeycomb or thin strips of lead and sometimes cadmium are used in photon path to reduce the amount of scatter. This enables lower doses to be given to the patient and also enhancing image quality. |  |
|  |  | Shielding NM-Detectors side and back shielding [150 kg] |  | 20000 | Detectors side and back shielding Reduction in image noise by natural radiation levels gives better image quality | N | The effectiveness of Gamma Rays shielding is a function of the density of the material. Any very dense material (eg Uranium) will shield against x-ray but may have other limitations. Lead offers the only realistic compromise between mass and volume. Detectors are sensitive to extraneous naturally occurring signals and need to be shielded from these signals to ensure only information from the region of interest are captured. |  |
|  |  | NM-Counterweights [50-100kg] |  |  | Counterbalance to movable heavy components to enable precise equipment positioning without patient involvement. | N | Some parts of gamma cameras are heavy but have to move in order to take optimum pictures, as it may be impossible/undesirable to move the patient. Conventionally counterbalances have been lead to be effective with the lowest volume. |  |
|  | Cadmium | NM Atomic ingredient of CdZnTe semiconductor radiation detector |  |  | 600 gm per 20cm x 20cm detector | N | Free cadmium is not used. It is bound in a stable (melting point = 1000º C) crystal Cd(Zn)Te. Similar to the Mercury in tooth fillings. Recommend recovery and return of discard. Small size of CdZnTe is necessary for advanced applications . |  |
|  | Cadmium(CdTe,CZT) | Radiation Semiconductor detector |  |  |  |  |  |  |
|  |  | measurement of radiation [3-15kg] (BvdE) |  | 300kg | X-Ray dose reduction for patient/image quality | ### N | Increase in the radiation s(th?)ickness subject. Reduction in performance (resolution, test time etc.) |  |
| Radiotherapy & Radiosurgery |  |  |  |  |  |  |  |  |
| Linear Accelerators | Lead | Shielding Shielding [200kg] |  | *3 (HS) | Reduction in unwanted ionizing radiation exposure | N | The effectiveness of shielding is a function of the density of the material. Any very dense material (e.g. Uranium) will shield against ionizing radiation but may have other limitations. Lead offers the only realistic compromise between mass and volume. |  |
|  |  | Counterweights [1-2 tons] (HS) |  |  | Counterbalance to enable equipment positioning to precisely target treatment point | N, Future design might make decreases | Linear Accelerators are heavy but have to move in order to provide optimum treatment. It is usually impossible/undesirable to move the patient. Conventionally counterbalances have been lead to be effective with the lowest volume. Counterweights are in principle re-used in next generation products without any effect on the environment |  |
|  | # Mercury | Check source thermometers |  |  | Measurement of air density for ionization chamber control | N | No adequate precision replacement available. |  |
| Multi-Source Stereotactic Radiosurgery (GammaKnife) | Lead | 804 kg |  |  | Reduction in unwanted ionizing radiation exposure | N | Multi-Source Stereotactic Radiosurgery (GammaKnife) |  |
| Brachytherapy Afterloader and Seedloader | # Lead | Lead acid batteries |  |  | power backup to preserve treatment history when mains power fails | Y 2012? | other technology is feasible, but requires mechanical and electronic design changes |  |
| Radiotherapy simulation |  |  |  |  |  |  |  |  |
| Simulators | Lead |  | Lead acid batteries |  | portability of system components (patient couch) | Y 2012? |  | other technology is feasible, but requires mechanical and electronic design changes |
| Ultrasound |  |  |  |  |  |  |  |  |
| All | Lead | Ultrasonic transducer crystals (0.5-3 g) |  | (80) (2) | Product reliability Improvement to image quality | N | No substitute substances due to the characteristics of ultrasound transmission/reception |  |
|  |  | Ultrasonic transducer crystal soldering |  |  | Product reliability and low cost | N | The lead-free solders require higher reflow temperatures which cause problems for performance of crystal |  |
|  |  | Probe/console Electrical connectors(e.g. PCB, connectors, wiring looms) |  |  | Product Reliability | Y, To be reviewed in 2012 | Lead solder is a well understood technology and consequently process reliability issues are to a large extent solved. Lead free solder is a relatively unknown process and so process reliability is still being developed. Other electrical Industries are converting to lead free and when the technology has stabilized, medical devices can make the change. |  |
| Electro-Medicine |  |  |  |  |  |  |  |  |
| Patient Monitoring | Lead | Solder on board |  |  | Reliability of product | Y To be reviewed in 2012 | In current condition, it is thought that it would take 3 to 5 years to secure reliability that solder excluding lead can be used for medical equipment. |  |
|  | Mercury | Display screens backlight |  |  |  | Possible, dependant on suppliers | Suppliers will have to deliver RoHS compliant screens |  |
| Life Support Systems | Lead | Solder on board |  |  | Reliability of product | Y To be reviewed in 2012 | In current condition, it is thought that it would take 3 to 5 years to secure reliability that solder excluding lead can be used for medical equipment. |  |
| ECG/EEG | Lead | Solder on board |  |  | Reliability of product | Y To be reviewed in 2012 | In current condition, it is thought that it would take 3 to 5 years to secure reliability that solder excluding lead can be used for medical equipment. |  |
| MCG (magnetoencephalography)/MEG (magnetocardiography) | Lead | Superconducting connections in superconducting SQUID sensors [< 1 g] |  | 20 g/year | Product reliability | Y? | Low-temperature lead-tin alloy bonding adheres reliably and has suitable mechanical properties to withstand mechanical stresses due to thermal contraction due to cooling to cryogenic temperatures. A very high reliability required for a multichannel instruments: hundreds of SQUIDs have to work simultaneously, repairing difficult and very costly. Only possibility for flip-chip connections. Alternative niobium wire bonding suffers from easy oxidization of surfaces and unfavourable mechanical properties, moreover requires heat treatment. All lead to reduced reliability. Marginal volume, yearly worldwide sales on the order of tens of systems. |  |
| Hearing Aids | Lead | Electrical connections (eg PCB, connectors, wiring looms) |  |  | Product reliability | Y To be reviewed in 2012 | Lead solder is a well understood technology and consequently process reliability issues are to a large extent solved. Lead free solder is a relatively unknown process and so process reliability is still being developed. Other electrical Industries are converting to lead free and when the technology has stabilized, medical devices can make the change. |  |
| General (for most Medical Devices) |  |  |  |  |  |  |  |  |
|  | Lead | Lenses [mg] |  |  | Optical Quality Colour Rendition X-ray reduction to safeguard image sensor | N | Lead in glass is essential for top optical performance, this is especially important for devices that use very small lenses like endoscopes. Suppliers will have to deliver RoHS compliant components |  |
|  |  | optical glasses for lenses and filters [mg] |  |  | refraction index, optical Quality Color Rendition X-ray reduction to safeguard image sensor | N | Lead in glass is essential for top optical performance, this is especially important for devices that use very small lenses like endoscopes. Suppliers will have to deliver RoHS compliant components |  |
|  |  | Electrical connections (eg PCB, connectors, wiring looms) [mg-3kg] |  |  | Product reliability (Patient safety, esp where the device controls the administration of energy and/or drugs to a patient) | Y, To be reviewed in 2012 | Lead solder is a well understood technology and consequently process reliability issues are to a large extent solved. Lead free solder is a relatively unknown process and so process reliability is still being developed. Other electrical Industries are converting to lead free and when the technology has stabilized, medical devices can make the change. (not before 2012) |  |
|  | Cadmium | optical glasses for lenses and filters [mg] |  |  | refraction index, optical Quality Color | N | Cadmium in glass is essential for top optical performance, this is especially important for devices that use very small lenses like endoscopes. Suppliers will have to deliver RoHS compliant components |  |
|  | Hexavalent Chromium | inorganic surfaces (FS) |  |  | image quality and reliability | N 2012 | Used to improve properties of surfaces e.g. media resistance in a clinical environment, sliding properties, corrosion, electromagnetic properties. |  |
|  | Alloy material lead |  |  |  | Improve material properties | N | E.g. in aluminium alloys. Used to improve cutting properties. |  |
| Other new technologies (e.g. molecular imaging) | Unknown | Tracers,,new technologies |  |  | Enabling gthe new technology |  | An exemption should be demanded if necessary. Flexibility for innovation for better healthcare should be maintained. |  |